## Slide 1: Tuesday 18.05
09:00 : Access in Pt4 for RB S34.
  Water flow low – increased.
11:40 : Recovered from PC fault and access. Beams back.
12:15 : RF sets up procedure for longitudinal emittance blow-up in
the SPS for high intensity operation.
13:30 : High intensity bunches injected
  Longitudinal emittance 0.6 eVs.
  Octupoles on.
  Ramp start 14:50:
  Emittance B1: ~ 6 mm in H and 4.3 mm in V. Emittance B2: ~ 8.2 mm in H and 5.3 mm in V.
  Flat top 15:45: B2 unstable  trim vert Q’ down by 6 units
  Emittance B1: ~ 10.7 um in H and 10.2 um in V. Emittance B2: ~ 11.2 um in H and 14.3 um in V.

## Slide 2: Tuesday 18.05
17:00 : Start collimator setup
  Observe loss of Landau damping in longitudinal plane (but beam still stable)
  17:30: Loss of Beam2 due to collimator motor position interlock in point 3. Controls issue.
  17:43: Loss of Beam1 due to interlock generated by ‘out of limit’ offset in centering between TCDQ and TCSG.
  Dump of Beam 1 was not clean with losses on the TCDS.
18:00 : Start ramp down
19:30 : Start preparation of physics fill over night
  21:40: Start injection of 6 x 2 10^2 bunches per beam & start of ramp
  22:11: several orbit correctors tripped due to feedback trim
  22:36: separation bumps collapsed
  23:34: lost more orbit correctors  Beam1 lost from orbit interlock @ TCDQ
  00:00: start rampdown

## Slide 3: ‘Dancing’ bunches at 3.5 TeV

## Slide 4: Beam 1 dump

## Slide 5: Wednesday 19.05
03:00 : new fill, ramp without FB.
07:30 : beam lost due to RCBX3.L1 trip
Presently ramping down
09:00 : Continue (Start) Collimation setup for High intensities.
(Possible) program for rest of week:
  Access UX45 for RF (UPS circuit breaker for RF crate)
  24h stop for Cryo de-icing from Thursday noon to Friday noon ?
  Try to finish collimation setup, collimator hierarchy
  RF measurements with 1e11 in the ramp
  Debunched beam issues with 1e11
  Transverse damper
  Physics fills….